## Slide 1: Hyperpolarized 13C MRI: Polarization & Imaging Physics
BioE241
02/25/2019

## Slide 2: Overview
DNP Physics
Polarizers
Imaging Approaches

## Slide 3: The Magnetic in MRI
Can think of 1H nuclei as dipoles
Align with or against the main magnetic field
So why the need for a strong magnet in MRI?

## Slide 4: The Magnetic in MRI
In the absence of a magnetic field, no preferred orientation

## Slide 5: Bulk Magnetization
What happens when spins are placed into a magnetic field?
  Zeeman splitting
  Precess about B0
    ω = γB0

## Slide 6: Bulk Magnetization
So why the need for a strong magnet in MRI?
  Earth’s magnetic field?
Boltzmann equilibrium!

## Slide 7: Boltzmann Distribution
Boltzmann equilibrium:
Net Magnetization (parallel - antiparallel)
Total # of spins

## Slide 8: SNR in MRI
Signal in an MR experiment is proportional to n, γ, and Pth.
Polarization is a function of γ, Bo and T:
At 3T and 37°C:
Pth(1H) ≈ 9 x 10-4 %
Pth(13C) ≈ 2.3 x 10-4 %

## Slide 9: Sensitivity in MRI
Water, fat have strong signal
13C, 15N, 31P inform about biochemistry
  Inherently poor SNR
Carbon is the backbone of life
  12C has no NMR signal
  13C abundance = 1.1%
Bottomley, Radiology 170(1): 1989

## Slide 10: Thermal Polarization
Polarization is a function of γ, Bo and T:
How can we increase polarization?
  Increase B0
  Decrease temperature

## Slide 11: Brute-Force Polarization
Force spins into high polarization at low temperature/high B0
Pros:
  Straightforward
Cons:
  Engineering challenge (3He/4He refrigerator)
  LONG buildup times

## Slide 12: Dynamic Nuclear Polarization
Exogenously increase 13C polarization via Dynamic Nuclear Polarization (DNP1)
Yields 20-30% 13C polarization
  >10,000 signal enhancement
1Ardenkjær-Larsen et al ., PNAS 2003; 100(18).

## Slide 13: DNP Hyperpolarization
t = 1s

## Slide 14: DNP Hyperpolarization
t = 1s
t = 65h

## Slide 15
DNP Requirements
Transfer polarization from e- to n by driving a forbidden ‘flip-flop’ two-quantum transition

## Slide 16
DNP Requirements
Transfer polarization from e- to n by driving a forbidden ‘flip-flop’ two-quantum transition
Requirements:
  Low temperature
  High field
  Free radical (unpaired electron)
    1:1000 ratio
  Forms a neat glass (no crystallization)
  MW source to transfer polarization from e- to nuclei

## Slide 17
DNP: Effect of Temperature and Field Strength
Polarizers operate at different field strengths and temperatures
  Hypersense: 3.35T/1.4K
  SPINlab: 3.35T/0.8K
      5.00T/0.8K
  Others operate at 7T-10.1T, 0.8K
In general, high field/lower temperature lead to higher polarization.
Why does this make a difference when P(e-) already 100%?

## Slide 18
DNP: Effect of Temperature and Field Strength
Comes down to interplay between T1(e-) and T1(n)
High field/low temperature predominantly leads to an increase in T1(n), increasing polarization
Doping with paramagnetic lanthanides (Gd, Ho) also increase polarization by preferentially reducing T1(e-)
T1(e)/T1(n)
'Leakage Factor’
Effects of nuclear
relaxation

## Slide 19
DNP: Effect of Radical
Many different stable radicals can be used
Choice of solvent may dictate which radical can be used (i.e. some radicals are hydrophobic, etc)
Narrow linewidth radicals (like trityl) are most effective for direct polarization of low-ɣ nuclei, like 13C
Radical concentration can also influence polarization
  Tradeoff between faster buildup times/increased T1(n) relaxation
    Data shown are at 7T

## Slide 20
Take 13C enriched compound
Add minute amount of stable unpaired electron
Cool to 1.4K at 3.35T
Irradiate with microwaves
3.35T magnet
Hypersense Polarizer

## Slide 21: Hypersense Polarizer

## Slide 22: Dissolution Procedure
The buffer is heated and pressurized
The sample space is pressurized
The sample is raised out of the liquid helium
The dissolution stick is lowered, docking with the sample holder
The solvent is injected, dissolving the sample, while preserving the enhanced polarization
~ 3.35T
~ 1.2K
Resulting sample:
High polarization
Physiologic temperature
Physiologic pH

## Slide 23: SPINlab Polarizer
Sterile fluid path
SpinLab + QC
QC Unit:
radical, oC, % polarization, pH

## Slide 24
GMP Chemicals
Low Bio-Burden Fluid Path
(FP)
Spinlab:
Polarize FP/Dissolve/terminal sterilization
Spin
PharmacyLab
Acquire mp-1H MRI and 13C MRI Data
inject
< 1min
Recon &
relate to other MR data
Filled FP
dispense
< 24hrs

## Slide 25
[113C]pyruvate
40,000 increase in MR signal at 3T1
1Ardenkjaer-Larsen et al. PNAS. 2003
2Kurhanewicz et al. Neoplasia, 2011
3Nelson et al., STM, 2013
Clinical MR Scanner
Volume 13C MRI in 10’s of secs3
Inject i.v. – Fast 13C MRI

## Slide 26: Hyperpolarized 13C has been able to…
Measure pH in tumors3
Measure cellular transport rates6
Give indicator to treatment response4
6Day et al. 2007
4Harris et al. 2009
Measure redox potential7
Measure blood flow and perfusion1,2
3Gallagher et al. 2008
2Ishii et al. 2007
Measure cardiac ischemia5
5Golman et al. 2008
1Mansson et al. 2006
7Keshari et al. 2011

## Slide 27
kPL
sec-1
0.016
0.004
0.008
Phase 2 Trial - 3D Dynamic HP 13C MRI in Patients Prior to and Following 6 months of Androgen Deprivation Therapy (ADT)
T2 wt. Image
ADC Image
Overlaid kpl Image
HP 13C spectral
Array
Pre-treatment
3 months post ADT (Lurpon + Casodex) + Doxcetaxel
Patient with Gleason 4+5 cancer and lymph node metasatses

## Slide 29
Imaging Probes

## Slide 30
Now that we have all this signal…
and all these probes…
how do we image them?

## Slide 31: Imaging Tradeoffs
SNR
Spatial
Resolution
Scan
Time
Tradeoff in conventional MRI between spatial resolution, temporal resolution, and SNR
With DNP, short scan times no longer hurt SNR
  Conversely, shorter scan time  improved SNR
  Biggest tradeoff with DNP is now between SNR and spatial resolution

## Slide 32: Hyperpolarized 13C Imaging: Pitfalls
Limited imaging time (T1 decay)
  30-60s for 13C species
Reduced γ
  42.56 MHz/T vs 10.71 MHz/T
RF Decay
  K-Space filtering1
  No averaging
  RF calibration
1Yen et al., MRM 62: 2009

## Slide 33: T1 Decay
S α exp(-T/T1)
T1 = 3s
T1 = 15s
T1 = 30s
T1 = 45s
T1 = 60s

## Slide 34: T1 Decay
T1 = 3s
T1 = 15s
T1 = 30s
T1 = 45s
T1 = 60s
S α exp(-T/T1)

## Slide 35: T1 Decay
4 main pathways for relaxation:
Highly dependent on local chemical environment
For quaternary carbons, primary relaxation pathway is chemical shift anisotropy:

## Slide 36: 1H vs 13C Imaging
Slice thickness and k-space trajectory modulated by γ:
Slice Thickness:
K-space Trajectory:
FOV:
Resolution:

## Slide 37: 1H vs 13C Imaging
Slice thickness and k-space trajectory modulated by γ:
Slice Thickness:
K-space Trajectory:
FOV:
Resolution:

## Slide 38: 1H vs 13C Imaging
Due to differences in γ, for the same trajectory 13C images have:
  4X larger FOV
  4X coarser resolution
  4X thicker slice
In other words, gradients have to work four times harder for 13C than for 1H

## Slide 39: RF Calibration
Endogenous 13C signal insufficient for power prescan
Crucial to calibrate power correctly
  Non-recoverable Mz
  Variable flip schedules
Can calibrate TG by:
  External 13C enriched syringe
  Bloch-Siegert B1 mapping1,2,3
1Schulte et al., NMR in Biomed 24: 2011.
2Lau et al., MRM 68: 2012.
3Sacolick et al., MRM 63: 2010.

## Slide 40: Constant Flip Angle
Received signal varies between excitations (can cause blurring)
Residual unused hyperpolarization after last excitation
S1
S2
S3
S4
S5
S6
θ
θ
θ
θ
θ
θ
MZ
MXY
Constant Flip Angle

## Slide 41: Constant Flip Angle
Zhao et al., JMR 113: 1996
`

## Slide 42
Flip angle is strictly increasing
Received signal constant when accounting for lost magnetization (Si = C)
Efficient usage of all polarization
θ1
S1
θ2
θ3
S2
S3
θ4
S4
θ5= 45°
S5
θ6 = 90°
S6
MZ
MXY
Variable/Progressive Flip Angle

## Slide 43: Variable/Progressive Flip Angle
Zhao et al., JMR 113: 1996

## Slide 44: Dynamic Imaging: Conventional Excitation
Lactate
13C1
Pyruvate
13C1
time
received signal
flip angle
2.5°
5°
10°
20°
Excess Pyruvate SNR
44
How to Detect HP Agents
Dynamic Imaging: Conventional Excitation

## Slide 45: Dynamic Imaging: Multiband Excitation
time
More lactate SNR for more time
Pyruvate SNR is still sufficient
Smaller pyruvate flip leaves more magnetization that can then become lactate
45
How to Detect HP Agents
Lactate
13C1
Pyruvate
13C1
flip angle
2.5°
5°
10°
20°
Dynamic Imaging: Multiband Excitation
received signal

## Slide 46: Multiband Variable Flip Angle in Primate Brain
Multiband variable flip angle: T1 = ∞, kPL = 0.07 1/s, 30 s acquisition time
Multiband constant flip angle: 20º lac, 4º pyr
No lactate observed
No pyruvate observed after 15 s
Park I, et. al. MRM, 2014.

## Slide 47: RF Coils
Need both 1H and 13C coils
  1H for localization, 13C for metabolism
Straightforward on a clinical scanner
  1H body coil built-in
If no 1H body coil, requires either:
  Swap between 1H and 13C coils
  Dual-tune 1H/13C coil for pre-clinical systems

## Slide 48: Imaging Approaches

## Slide 49: k-Space: A Brief Overview
Want to turn NMR signal into a spatially-resolved image
Accomplish by encoding magnetization with gradients in k-space (frequency domain):
For HP substrates, need to encode through 5 dimensions: space (X/Y/Z), time, and frequency

## Slide 50: Hyperpolarized Imaging
Metabolic Substrates
Inert Substrates

## Slide 51: Hyperpolarized Imaging: Inert Substrates
Only one resonance present
  Urea, HP001, t-butanol1
Can use conventional excitation and a wide array of fast imaging techniques
  Spiral
  EPI
  bSSFP1,2
  GRASE3
1von Morze et al., MRM 72: 2014.
2Reed et al., IEEE TMI 33(2): 2014.
3Sukumar, ISMRM 2012, #4292.

## Slide 52: Hyperpolarized Imaging: SSFP with Inert Substrates
1von Morze et al., MRM 72: 2014.
2Reed et al., IEEE TMI 33(2): 2014.
3Sukumar, ISMRM 2012, #4292.
Pros:
  Highest SNR efficiency
  Most beneficial with long T2 substrates
  Highest resolution
    1.0mm3 isotropic
Cons
  Banding artifacts
    Stopbands at ±1/TR
  Difficult with metabolic substrates
13C, 15N Urea
1 volume every 2.4s

## Slide 53: Hyperpolarized Imaging: Metabolically Active Substrates
Non-equilibrium magnetization:
Metabolism, RF and T1 Decay
Golman et al., Cancer Res 66(22): 2006
Rowland et al., Curr Pharm Biotechno 11(6): 2011

## Slide 54: Dealing with Off-Resonance
Off-resonance leads to artifacts
  Position shifts in Cartesian
  Blurring/Distortions in non-Cartesian
    Worse with long readouts
Two major sources of off-resonance
  B0 heterogeneity (shim quality, susceptibility)
  Chemical Shift of species
Two main classes of acquisitions:
  Spectroscopic imaging
  Frequency-specific imaging

## Slide 55: Spectroscopic Imaging Techniques
FIDCSI1,2
1Golman et al., PNAS 2006
2Kohler et al., MRM 2007
EPSI3,4
3Cunningham et al., JMR 2007
4Larson et al., MRM 2011
Spiral IDEAL5,6
5Wiesinger et al., MRM 2012
6Gordon et al., ISMRM 2012
Spiral CSI7
7Mayer et al., MRM 2006
k-t Spiral8
8Gordon et al., MRM 2013

## Slide 56: CSI: Chemical Shift Imaging1,2
Phase encode on both axes
Pros:
  High spectral resolution
  High spectral bandwidth
Cons: quite slow!
  Scan time = TR x Nx x Ny
  RF inefficient!
    # RF = Nx x Ny
1Kohler et al., MRM 58: 2007.
2Hurd et al., MRM 63:2010.

## Slide 57: Rapid Spectroscopic Techniques
Spiral CSI, EPSI, etc.
Spectral encoding accomplished by repeating spatial gradients shifted in time by ΔTE
  SW = 1/ΔTE
  Resolution = SW/# echoes
Spatial/spectral interleaves required to fully sample kx-ky-kf space

## Slide 58: Rapid Spectroscopic Techniques
Pros:
  More RF efficient than CSI
  Faster then CSI
  Scan time = TR x NPE
Downsides:
  Low spectral bandwidth (1/ΔTE)
    ΔTEEPSI ≈ 2-4ms1
    ΔTESpiral CSI = 3.6ms2
1Yen et al., MRM 62: 2009
2Mayer et al., MRM 65:2011
Adalsteinsson et al., MRM 1998
Cunningham et al., JMR 2007

## Slide 59: Rapid Spectroscopic Techniques
Pros:
  More RF efficient than CSI
  Faster then CSI
  Scan time = TR x NPE
Downsides:
  Low spectral bandwidth (1/ΔTE)
    ΔTEEPSI ≈ 2-4ms1
    ΔTESpiral CSI = 3.6ms2
Tradeoff between spatial resolution and spectral bandwidth3 due to ɣ(13C)
1Yen et al., MRM 62: 2009
2Mayer et al., MRM 65:2011
3Jiang et al., DOI: 10.1002/mrm.25577

## Slide 60: Model-Based Reconstructions
Hyperpolarized 13C spectra are sparse
  A-priori knowledge of resonances
Can use model-based approaches to minimize imaging requirements1,2,3
  Can use same acquisition as EPSI or spiral CSI, but:
  No longer need to have sufficient spectral BW
    Allow spectral aliasing, flexible choice of ΔTE
  Arbitrary number of echoes (# echoes > # freq + 1)
  Can incorporate field map, multiple coils (parallel imaging), compressing sensing, etc. into reconstruction
3Gordon et al., MRM 71(4): 2014
1Reeder et al., JMRI 26: 2007
2Wiesinger et al., MRM 68(1): 2012

## Slide 61: Signal Model in k-space1,2
With a known (or ignored) field map, this becomes a linear problem
Solve for ρm with a least-squares approach point-by-point in k-space
Signal at k-space point k sampled at echo n
Known phase from chemical shift
Unknown spin density
1Brodsky et al., MRM 59: 2008
2Wiesinger et al., MRM 68(1): 2012

## Slide 62
RF
Spiral Readouts
TE1
TE2
TE3
TE4
Spectral Images
Pyruvate
Lactate
Pyruvate-Hydrate
Alanine
Matrix Inversion
FFT
Spectral
K-space
1Brodsky et al. MRM 59(5): 2008
2Wiesinger et al., MRM 68(1): 2012
Model-Based Reconstruction

## Slide 63: Spectral Selectivity
Use spectrally selective RF pulses to control flip angles for different compounds
Conventional RF pulses (i.e. sinc) can only be either spatially or spectrally selective

## Slide 64: Spectral-Spatial RF Pulses1,2,3
Spectral-spatial pulses are selective in both frequency and space
Add oscillating gradient for additional spatial selectivity
  Subpulses control spatial selectivity
  # of subpulses and their separation control spectral selectivity
Useful in vivo where spatial selectivity is important
1Lau et al., NMR in Biomedicine 2011
2Larson et al., JMR 2008
3Kerr et al., ISMRM 2008: #226

## Slide 65: Spectral-Spatial RF Pulse Design
Larson et al., JMR 2008
Kerr et al., ISMRM 2008: #226

## Slide 66: Metabolite-Specific Imaging
Use the RF pulse for spectral selection1
Just like with inert substrates, now amenable to fast imaging techniques
  Spiral2
  EPI3,4
  GRASE5
Want to keep the readout ≤ T2* to minimize blurring
3Cunningham et al., JMR 193: 2008.
4Gordon et al., MRM, 2017.
1Lau et al., NMR in Biomed 24: 2011.
2Lau et al., MRM 64: 2010.
5Sukumar, ISMRM 2012, #4292.
f = lac, ala…
m,n = 8,12,16

## Slide 67: Metabolite-Specific Imaging
Pros
  Fast!
    Scan time = TR x # slices
  Can use any fast imaging approach
Cons
  Long RF pulses (15-20ms)
  Sensitive to off-resonance
Question: Could we use SSFP here?

## Slide 68: Comparison of Acquisition Schemes1
Ultimately have to decide what’s the most important to you:
  Temporal resolution?
  Spectral bandwidth/resolution?
  SNR efficiency?
  Robustness?

## Slide 69: Comparison of Acquisition Schemes1
1Durst et al., NMR in Biomedicine,28(6) : 2015.

## Slide 70: Comparison of Acquisition Schemes1
1Durst et al., NMR in Biomedicine,28(6) : 2015.

## Slide 71: Sequence Tradeoffs
 | Pros | Cons
Metabolite-Specific Imaging | Speed High SNR efficiency | Sensitive to off-resonance Metabolites must have sufficient chemical shift (~>2ppm apart)
CSI | High spectral bandwidth High spectral resolution | Slowest Most RF inefficient
MRSI (EPSI, Spiral CSI) | Robust to off-resonance Acquire entire spectrum No prior knowledge req’d | Slow RF inefficient Limited spectral bandwidth
Model-Based Approaches (IDEAL) | Fast High SNR efficiency Few RF pulses required | Must have prior knowledge of Δf and # of peaks Optimal ΔTE may not be obtainable

## Slide 72: Summary & Conclusions
Multitude of options for acquiring hyperpolarized data
Think about your substrate-specific considerations
Each sequence has strengths/weaknesses
Good review/overview articles:
  Durst - Comparison of acquisition schemes for hyperpolarised 13C imaging
  Gordon & Larson - Pulse Sequences for Hyperpolarized MRS
  Niedbalski - NMR Spectroscopy Unchained: Attaining the Highest Signal Enhancements in Dissolution Dynamic Nuclear Polarization

## Slide 73: Sequences available at UCSF: 3T GE (Clinical Scanner)
1D spectroscopy (global/slice-selective)
2D CSI
Metabolite-Specific Imaging
  EPI
  Spiral
Model-Based Approaches
  Spiral
EPSI (2D/3D)
SSFP

## Slide 74: Sequences available at UCSF: 3T Bruker (Pre-Clinical Scanner)
1D spectroscopy (global/slice-selective)
2D CSI
2D EPSI
Metabolite-Specific Imaging
  SPSP + EPI

## Slide 75: Sequences available at UCSF: 14T Varian (Pre-Clinical Scanner)
1D spectroscopy (global/slice-selective)
2D/3D CSI
2D EPDI
Metabolite-Specific Imaging
  SPSP + EPI
  With diffusion weighting (optional)
GRASE
bSSFP